## Sheet1
TRON Programming Contest 2025
BOM REV : 0
| Sr.no | Ref. | Part Description | Mfg. Part Details | Qty. | Unit Cost | Total Cost | Purpose | Functionality
| 1 | Flyback Diode | 1N4148 | Link | Pack of 10 | Rs. 29 | Rs. 29 | Protects MCU/GPIO from inductive spikes | Clamps voltage spike when vibration motor is switched off
| 2 | Dry In-Ear EEG Electrodes | Disposable/Reusable Longer Spike dry EEG Electrode TDE 210 – 5mm | Link | 3 | Rs. 400 | Rs. 1566 | Acquire EEG signals from ear canal | 5mm spike comb electrodes for microvolt-level EEG capture without gel
| 3 | Coin-Type Vibration Motor | 3 V, 10 mm | Link | 2 | Rs. 22 | Rs. 44 | Provide discreet haptic alerts | 10 mm eccentric-mass motor, draws ~100 mA peak, fits inside earbud shell
| 4 | Logic-Level MOSFET | 2N7002 | Link | 4 | Rs. 3 | Rs. 12 | Drive haptic motor safely | 60 V, 300 mA N-channel SOT-23 MOSFET, gate driven by 3.3 V GPIO
| 5 | MEMS Microphone Module | INMP441 I²S | Link | 2 | Rs. 198 | Rs. 396 | Capture voice commands & ambient noise | Digital I²S microphone, omnidirectional, low noise floor
| 6 | ADC Module | ADS1263 32-bit with PGA breakout | Link | 1 | Rs. 2802 | Rs. 2702 (Rs.100 Discount) | High-resolution EEG signal acquisition | 32-bit ΔΣ ADC, 2-channel differential, PGA, SPI interface - requires additional DRL circuit
| 7 | DRL Op-Amp | LM358 | Link | 1 | Rs. 10 | Rs. 10 | Driven-Right-Leg amplifier for common-mode rejection | Low-noise op-amp to create DRL feedback, reduces 50/60 Hz interference
| 8 | Bias Resistors | 100 M Ohms | Link | 10 | Rs. 108.32 | Rs. 1007.38 | Electrode bias to VREF/2 | High-value resistors for biopotential input bias, sets common-mode when electrodes detach
| 9 | Anti-Alias Capacitors | 10nF | Link | Pack of 10 | Rs. 36 | Rs. 36 | Input filtering & protection | Forms anti-alias filter with bias resistors, ~160 Hz cutoff
| 10 | Li-Po Battery | 3.7 V, 450 mAh | Link | 2 | Rs. 277 | Rs. 554 | Power earbud electronics | Protected pouch cell, ~22 h runtime with duty-cycled sensing/inference
| 11 | Li-Po Charger IC | MCP73831T-2ACI-OT | Link | 2 | Rs. 89 | Rs. 178 | Battery charging control | Single-cell Li-Po charge management, up to 500 mA, SOT-23-5 package
| 12 | Fuel Gauge | MAX17048 | Link | 1(Robocraze) + 1(Evelta) | Rs. 194 & Rs. 247 | Rs. 441 | State-of-charge monitoring | Coulomb counter + voltage/temperature monitoring, reports SoC/SoH via I²C
| 13 | Sillicone Wire | 30 AWG Sillicone Wire Kit | Link | 1 | Rs. 449 | Rs. 449 | Provide ultra-fine, highly flexible, silicone-insulated leads for each EEG electrode connection (Left +, Left –, DRL/common, Right +, Right –) | Multi-stranded 30 AWG conductors minimize signal attenuation and microphonic noise.Silicone insulation remains supple and durable inside the earbud shell.
| 14 | Breadboard | GL12 840 Points Solderless Breadboard | Link | 1 | Rs. 54 | Rs. 83.5 | Connections | Connections
| 15 | Jumper Wires | Jumper Wire Set - M2M, M2F, F2F (40 pcs each) | Link | 1 | Rs. 129 | Rs. 158.5 | Connections | Connections
| 16 | Instrumentation Amplifier | ADS623 | Link | 1 | Rs. 561 | Rs. 648 | To amplify EEG signals with precision | The AD623 is specifically designed to amplify small differential signals while rejecting common-mode noise. It offers high input impedance and a high CMRR, both of which are essential for bio-signal measurement.
| 17 | Low Dropout Voltage Regulator (LDO) | AMS1117 | Link | 1 | Rs. 33 | Rs. 33 | Step down voltage | Power profile enhancement and user safety
| 18 | Zener Diode | 1N4728A | Link | 10 | Rs. 10 | Rs. 193 | Protects against ESD Discharge | Terminates sudden ESD spikes
GROSS : Rs. 8540.38
***(The gross total also includes delivery charges, which are directly included in the “Total Cost” column)